Verification of Analytical Balance Weights Policy

(NC WW/GW LC Branch 04/20/2020)

The purpose of this policy is to expand upon and explain the requirements described in 15A NCAC 02H .0805 (a) (7) (J), which states, in part: The analytical balance shall be checked with one ASTM Type 1, Class 1 or 2, or equivalent standard weight each day used. These weights shall be verified every five years.

Verification may be accomplished by:

- Sending laboratory weights to an ISO/IEC 17025:2017 compliant organization that can provide traceability showing compliance with maximum tolerances required in the ASTM specifications below, or;

- Checking laboratory weights against certified reference weights (i.e., weights that have been recertified as in 1. above) and found to be within ASTM Class 1 or Class 2 tolerances (see table below).  The balance service technician may provide this service if a balance with the proper resolution (i.e., 0.000001 gram) is available.

Note: Although some manufacturers will assign a one-year calibration due date, the NC WW/GW LC Branch considers 5 years to be an acceptable calibration interval.  Laboratories may want to request that a 5-year calibration due date be stated on the traceability documentation when having weights verified or when ordering new weights.

Documentation of weight verifications or recertification must be maintained for 5 years. If the condition of any weight is in question at any time due to damage (e.g., corrosion, nicks, scratching, etc.), that weight must be re-verified as described above and in accordance with 15A NCAC 02H .0805 (a) (6).

Maximum Tolerances (Ref. ASTM E 617-18, 2018)

| Denomination | Maximum Tolerance for ASTM Class 1 Weights (± mg) | Maximum Tolerance for ASTM Class 2 Weights (± mg) |
| --- | --- | --- |
| 1 kg | 2.5 | 5.0 |
| 500 g | 1.2 | 2.5 |
| 300 g | 0.75 | 1.5 |
| 200 g | 0.50 | 1.0 |
| 100 g | 0.25 | 0.50 |
| 50 g | 0.12 | 0.25 |
| 30 g | 0.074 | 0.15 |
| 20 g | 0.074 | 0.10 |
| 10 g | 0.050 | 0.074 |
| 5 g | 0.034 | 0.054 |
| 3 g | 0.034 | 0.054 |
| 2 g | 0.034 | 0.054 |
| 1 g | 0.034 | 0.054 |
| 500 mg | 0.010 | 0.025 |
| 300 mg | 0.010 | 0.025 |
| 200 mg | 0.010 | 0.025 |
| 100 mg | 0.010 | 0.025 |
| 50 mg | 0.010 | 0.014 |
| 30 mg | 0.010 | 0.014 |
| 20 mg | 0.010 | 0.014 |
| 10 mg | 0.010 | 0.014 |
| 5 mg | 0.010 | 0.014 |
| 3 mg | 0.010 | 0.014 |
| 2 mg | 0.010 | 0.014 |
| 1 mg | 0.010 | 0.014 |
| 0.5 mg | 0.010 | 0.014 |
| 0.3 mg | 0.010 | 0.014 |
| 0.2 mg | 0.010 | 0.014 |
| 0.1 mg | 0.010 | n/a |
| 0.05 mg | 0.010 | n/a |