May 23, 2019

# FECAL COLIFORM REPORTING

The following criteria are to be used in obtaining and reporting fecal coliform data:

Standard Methods suggests analyzing samples by filtering three different volumes (diluted or undiluted) depending on the bacterial density. Each laboratory must filter multiple dilutions of the sample in order to obtain plates containing 20 to 60 fecal coliform colonies. It is recommended that sampling containers of at least 250 ml be used in order to collect sufficient volume to meet  method criteria.  The  requirements for calculating values are given in Standard Methods 9222 D-2006, and EPA Microbiological Methods for Monitoring the Environment, Water, and Wastes, EPA/600/8-  78/017. 1978. US EPA. Page 124. The following is a compilation of these requirements to be used in calculating the fecal coliform count   per 100 ml    of sample for compliance with NC WW/GW Laboratory Certification.

# ALL RESULTS MUST BE REPORTED IN WHOLE NUMBERS.

- Countable Membranes with 20-60 Blue Colonies: Calculate the fecal coliform results from membrane filters within the ideal counting range of 20-60 blue colonies using the general formula:

Number of colonies counted x 100 = Fecal coliform colonies per 100 ml volume of sample filtered in ml

If more than one filter (including a 100 ml sample volume) has a count in the acceptable range, calculate the values in counts/100 ml for each filter in the countable range and average.

- Countable Membranes with less than 20 Blue Colonies: If all counts are below the lower limit (20) of the ideal counting range:

  - Select the count most nearly acceptable and compute the count  using  the  general  formula.  Report  the  count  as  an  Estimated  Count per 100 ml: or

  - Total the counts on all filters and report as number per 100 ml. For example, if 50, 25, and 10 ml portions were examined, and counts were 15, 6, and 0 coliform colonies respectively, calculate results as follows and report the count as 25 colonies per 100 ml.

(15 + 6 + 0) counts x 100 = 25 colonies per 100 ml

50 + 25 + 10 ml

- Membranes with No Colonies:  If counts from all filters are zero, report the count for the fecal coliform as a less than (<) value. Calculate  the number of colonies per 100  ml  that  would  have  been  reported  if  there  had  been  one  colony  on  the  filter  representing  the  largest filtration volume. For example, sample volumes of 25, 10 and 2 ml produced colony counts of 0, 0 and 0, respectively. The count  would be reported as <4 colonies per 100 ml. If a 100 ml sample was analyzed report as < 1/100 ml.

<1 counts x 100 = <4 colonies per 100 ml 25 ml

- Countable Membranes with more than 60 Colonies: If all filter counts are above the upper limit (60), but countable, calculate the count from the smallest volume filtered and report as a greater than (>) value. For example, if 10, 5, and 1  ml portions  of  samples  were  examined, and counts were TNTC, 310, and 95 coliform colonies respectively, calculate results as follows and report the count as >9500 colonies/100 ml.

>95 counts x 100 = >9500 colonies per 100 ml 1 ml

- Uncountable Membranes: For uncountable filters with more  than 60 colonies or "Too numerous  to Count"  (TNTC),  use 60  colonies as  the basis of calculation with the smallest filtration volume and report as a greater than value. For example, sample volumes of 10, 1.0 and

0.1 ml all produced too many colonies to show separated colonies and the laboratory bench sheet showed TNTC. The count would be reported as >60,000 colonies per 100 ml.

>60 counts x 100 = >60,000 colonies per 100 ml

0.1 ml

- If the Filters for a sample have counts of both >60 and <20, but none in the 20-60 range: Use all countable filters and calculate as in

(2) (b) above.

- Anomalies: The above requirements are to be used except when an abnormality  occurs  in  the  analysis  of  a  sample.  When  abnormalities occur, analysts must use their best judgment in selecting the proper value to report.

- Reporting Data for Multiple Samples in a Single Day or Duplicate Samples: When only one dilution is duplicated, that would be included in the sample calculation and not reported as separate result. When the entire dilution scheme for a sample is duplicated or if more than one sample is collected in a single day, all results must be reported. The DWR Water Quality Permitting Section has directed that those results may be reported discretely, or a as a geometric mean. If reported discretely, the highest value would go in the daily cell and the rest would go in the comment section. If the geometric mean is reported, the comment section must indicate the reported value is a geometric mean of either a duplicate sample or of multiple samples collected on that day.

Revised 05/23/2019