# Luminescence DO (LDO) Daily Check

Although many manufacturers of LDO probes indicate that the probes require calibration only every 6 months or so, NC WW/GW LC requires that an LDO probe must be calibrated or have the calibration verified each day of use. Below is a procedure for verifying the calibration of an LDO probe.

- Place probe in a plastic bag with a wet sponge or a BOD bottle partially filled with water

- Make sure the bag is effectively sealed (zip, rubber band, or twist tie)

- Allow appropriate instrument warm up time

- Read D.O. and temperature

- Check the reading vs. the Solubility Table below and apply appropriate atmospheric (barometric) pressure or altitude correction factor

| Temp. °C | D.O. mg/L |  | Temp. °C | D.O. mg/L |
| --- | --- | --- | --- | --- |
| 4 | 13.11 |  | 19.5 | 9.18 |
| 4.5 | 12.94 |  | 20 | 9.09 |
| 5 | 12.77 |  | 20.5 | 9.00 |
| 5.5 | 12.61 |  | 21 | 8.92 |
| 6 | 12.45 |  | 21.5 | 8.83 |
| 6.5 | 12.30 |  | 22 | 8.74 |
| 7 | 12.14 |  | 22.5 | 8.66 |
| 7.5 | 11.99 |  | 23 | 8.58 |
| 8 | 11.84 |  | 23.5 | 8.50 |
| 8.5 | 11.70 |  | 24 | 8.42 |
| 9 | 11.56 |  | 24.5 | 8.34 |
| 9.5 | 11.42 |  | 25 | 8.26 |
| 10 | 11.29 |  | 25.5 | 8.18 |
| 10.5 | 11.16 |  | 26 | 8.11 |
| 11 | 11.03 |  | 26.5 | 8.04 |
| 11.5 | 10.90 |  | 27 | 7.97 |
| 12 | 10.78 |  | 27.5 | 7.90 |
| 12.5 | 10.66 |  | 28 | 7.83 |
| 13 | 10.54 |  | 28.5 | 7.76 |
| 13.5 | 10.42 |  | 29 | 7.69 |
| 14 | 10.31 |  | 29.5 | 7.62 |
| 14.5 | 10.20 |  | 30 | 7.56 |
| 15 | 10.08 |  | 30.5 | 7.50 |
| 15.5 | 9.98 |  | 31 | 7.43 |
| 16 | 9.87 |  | 31.5 | 7.37 |
| 16.5 | 9.77 |  | 32 | 7.31 |
| 17 | 9.67 |  | 32.5 | 7.24 |
| 17.5 | 9.57 |  | 33 | 7.18 |
| 18 | 9.47 |  | 33.5 | 7.12 |
| 18.5 | 9.38 |  | 34 | 7.07 |
| 19 | 9.28 |  | 34.5 | 7.01 |

Calculated D.O. value must verify meter reading within ± 0.5 mg/L (do NOT calculate and apply a correction factor to calculated D.O.).

| Atmospheric Pressure mm Hg | Equivalent Altitude Ft. | Correction Factor |
| --- | --- | --- |
| 760 | 0 | 1.00 |
| 752 | 278 | .99 |
| 745 | 558 | .98 |
| 737 | 841 | .97 |
| 730 | 1126 | .96 |
| 722 | 1413 | .95 |
| 714 | 1703 | .94 |
| 707 | 1995 | .93 |
| 699 | 2290 | .92 |
| 692 | 2587 | .91 |
| 684 | 2887 | .90 |
| 676 | 3190 | .89 |
| 669 | 3496 | .88 |
| 661 | 3804 | .87 |
| 654 | 4115 | .86 |
| 646 | 4430 | .85 |
| 638 | 4747 | .84 |
| 631 | 5067 | .83 |
| 623 | 5391 | .82 |
| 616 | 5717 | .81 |
| 608 | 6047 | .80 |
| 600 | 6381 | .79 |
| 593 | 6717 | .78 |

Ref: YSI Model 5000/5100 DO Meter Manual. Slight variations in DO, pressure, and/or altitude may be found in other manuals.

Example: If ambient temperature is 21°C and elevation is approximately 1126 ft, the theoretical DO would be:

8.92 X 0.96 = 8.56 mg/L

or, If ambient temperature is 21°C and the atmospheric (barometric) pressure is 745 mm Hg, the theoretical DO would be:

8.92 X 0.98 = 8.74 mg/L