Mechanical Volumetric Liquid-Dispensing Devices Calibration (NC WW/GW LC Policy 10/02/2013)

Mechanical volumetric liquid-dispensing devices (e.g., fixed and adjustable auto-pipettors, bottle-top dispensers, etc.), used for critical measurements, must be calibrated at least twice per year, approximately six months apart. Each liquid-dispensing device must meet the manufacturer’s statement of accuracy. For variable volume devices used at more than one setting, check the accuracy at the maximum, middle and minimum values. Testing at more than three volumes is optional. When a device capable of variable settings is dedicated to dispense a single specific volume, calibration is required at that setting only.