# Method Blank Analysis Requirement for Suspended, Dissolved and Total Residue Policy

(NC WW/GW LCB 7/19/2021)

The analysis of method blanks for the Residue methods shall be performed as follows:

Total Suspended Residue: Using the same containers and glassware normally in contact with samples, put 30 ml of DI water through the sample filter and proceed through the entire analytical process. Acceptance criterion is <0.5 mg weight gain.

Total Dissolved Residue: Using the same containers and glassware normally in contact with samples, put

30 ml of DI water through the sample filter and proceed through the entire analytical process. Acceptance criterion is <0.5 mg weight gain.

Total Residue: No water is used. Take a prepared dish through the entire analytical process. Acceptance

criterion is <0.5 mg weight gain.