| LABORATORY NAME: |  |  | CERT #: |  |
| --- | --- | --- | --- | --- |
| PRIMARY ANALYST: |  |  | DATE: |  |
| NAME OF PERSON COMPLETING CHECKLIST (PRINT): |  |  |  |  |
| SIGNATURE OF PERSON COMPLETING CHECKLIST: |  |  |  |  |

Parameter: Volatile Petroleum Hydrocarbons (VPH)

Method: Massachusetts Department of Environmental Protection VPH, February 2018 Rev. 2.1 (Aqueous & Non-Aqueous)

|  | Purge & Trap System - Model: |  |
| --- | --- | --- |
|  | Gas Chromatograph - Model: |  |
|  | PID & FID – Model: |  |
|  | Capillary GC Column: | 105 m x 0.53 mm 1.0., 3.0 µm Restek Rtx- 502.2 or equivalent (Required) |

| Equipment: |  |  | Type of Trap: |  |
| --- | --- | --- | --- | --- |
|  | Syringes |  |  |  |
|  | Syringe valves |  |  |  |
|  | Micro syringes |  |  | Equivalent to the following? (Y/N): |
|  | Analytical Balance |  | 400 mg of Carbopack B 7.6 cm Carbopack B and 1.3 cm Carbosieve S-W (opt.) 7.7 cm Carbopack C and 1.2 cm Carbopack B (opt.) |  |
|  | 103° - 105°C oven |  | Reagents: |  |
|  | VOC Vials, 40-mL (60-mL optional for soil) |  |  | Methanol, purge-and-trap grade |
|  |  |  |  | Hydrochloric acid: HCl, 1 + 1 |
| Standards: |  |  |  | Volatile-free Sand |
|  | Calibration Standards |  |  | Ascorbic Acid |
|  |  |  |  | Reagent Water |
|  | Surrogate |  |  |  |

| PLEASE COMPLETE CHECKLIST IN INDELIBLE INK Please mark Y, N or NA in the column labeled LAB to indicate the common lab practice and in the column labeled SOP to indicate whether it is addressed in the SOP. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | GENERAL | LAB | SOP | EXPLANATION |  |
|  | Is the SOP reviewed at least every 2 years? What is the most recent review/revision date of the SOP? [15A NCAC 02H .0805 (a) (7)] Date: |  |  | Quality assurance, quality control, and Standard Operating Procedure documentation shall indicate the effective date of the document and be reviewed every two years and updated if changes in procedures are made. Verify proper method reference. During review notate deviations from the approved method and SOP. |  |
|  | Are all review/revision dates and procedural edits tracked and documented? [15A NCAC 02H .0805 (a) (7)] |  |  | Each laboratory shall have a formal process to track and document review dates and any revisions made in all quality assurance, quality control and SOP documents. |  |
|  | Is there North Carolina data available for review? |  |  |  |  |
|  | PRESERVATION and STORAGE | LAB | SOP | EXPLANATION |  |
|  | Are aqueous samples collected in triplicate without headspace in contaminant-free 40 mL glass VOC vials with Teflon lined septa screw caps? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.1.1 (a)] |  |  | No air bubbles allowed in 40 mL vial. |  |
|  | Are aqueous samples preserved to pH ≤2 S.U. with 1:1 HCl at the time of collection [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.1.1 (a)] |  |  | Note: All preservation requirements in this checklist assume analysis with an ambient purge temperature. |  |
|  | Is the weight of the non-aqueous sample vials with preservative determined prior to sample collection? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.3.1] |  |  | Weigh the sample vial to 0.1 g on a top-loading balance and determine the weight of the soil/sediment sample; this determination requires knowledge of the empty/tared weight of the sample vial and volume/weight of methanol preservative that was added to the sample vial. |  |
|  | Are non-aqueous samples collected in duplicate in 60-mL or 40-mL VOC vials with Teflon-lined septa screw caps? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.2.4] |  |  | Samples for VPH analysis should be collected in duplicate 60-mL or 40-mL VOC vials with Teflon-lined septa screw caps. An additional sample of the soil/sediment must also be obtained (without methanol) to allow for a determination of moisture content and VPH dry weight correction factors. |  |
|  | Are non-aqueous samples completely covered with methanol preservative upon collection? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 8.2.2 and 8.2.3] |  |  | Methanol preservation of soil/sediment samples is mandatory. Methanol (purge-and-trap grade) must be added to the sample vial before or immediately after sample collection. In all cases, the soil/sediment sample in the vial must be completely covered by methanol. |  |
|  | Is an additional sample of the soil/sediment obtained without methanol to allow for a determination of moisture content? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.2.4] |  |  |  |  |
|  | Are samples cooled to 0-6°C immediately after collection? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.1.1 (a) & 8.2.1] |  |  |  |  |
|  | Are aqueous samples analyzed within 14 days of collection? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.1.5] |  |  |  |  |
|  | Are soil/sediment samples analyzed within 28 days of collection? [MADEP VPH Rev. 2.1 (Feb 2018) Section 8.2.8] |  |  |  |  |
|  | What corrective action is taken if samples do not meet the thermal and/or chemical preservation requirements? [15A NCAC 02H .0805 (a) (7) (M)] |  |  |  |  |
|  | PROCEDURE – Reagents and Standards | LAB | SOP | EXPLANATION |  |
|  | Are stock standards stored with minimal headspace at -10°C to -20°C and protected from light? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.2] |  |  | Store, with minimal headspace, at -10°C to -20°C and protect from light. |  |
|  | Are stock standards disposed of within 6 months of opening or by the manufacturer’s expiration date, whichever comes first? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.2] |  |  | Stock standard solutions must be replaced after 6 months, or sooner if comparison with check standards indicates a problem. |  |
|  | How are primary dilution standards stored? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.3] Answer: |  |  | These standards should be stored with minimal headspace, at -10°C to -20°, and should be checked frequently for signs of degradation or evaporation. |  |
|  | How often are the primary dilution standards replaced? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.3] Answer: |  |  | The primary dilution standards should be replaced at least monthly. |  |
|  | How long are aqueous calibration standards used after preparation? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.4] Answer: |  |  | Aqueous standards are not stable and should be discarded after one hour. |  |
|  | PROCEDURE – Instrument Calibration | LAB | SOP | EXPLANATION |  |
|  | Are a minimum of 5 calibration standards analyzed with concentrations evenly dispersed over the full working range and with the lowest calibration point corresponding to the reporting limit (RL)? [MADEP VPH Rev. 2.1 (Feb 2018) Section 7.4] |  |  | At a minimum, five different concentrations are required for a valid calibration curve. The calibration concentrations must be evenly dispersed over the full working range of the detector with the lowest calibration point corresponding to the RL. |  |
|  | How are retention time windows established over the course of a 72-hour period for each GC and any time a new column is installed? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.3] Answer: |  |  |  |  |
|  | Are the collective aliphatic/aromatic hydrocarbon ranges defined per the method? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.3.6] |  |  | The Rt window of the C5-C8 aliphatic hydrocarbons is defined as beginning 0.1 minutes before the elution of n-pentane and ending 0.01 minutes before the elution of nonane. The C9-C12 aliphatic hydrocarbon range begins 0.01 minutes before the elution of nonane; therefore there is no overlap of the two ranges and the nonane peak is only included in the C9-C12 aliphatic hydrocarbon range. The C9-C12 aliphatic hydrocarbon range ends 0.1 minutes before the elution of naphthalene. The Rt window for the C9-C10 aromatic hydrocarbons is defined as beginning 0.1 minutes after the elution of o-xylene and ending 0.1 before the elution of naphthalene. |  |
|  | Are the collective aliphatic/aromatic hydrocarbon ranges calibrated using Calibration Factors (CFs) according to the method? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.4.1.2] If CFs are not used, skip to #24 |  |  | CFs are calculated for C5-C8 aliphatic hydrocarbons and C9-C12 aliphatic hydrocarbons based upon a correlation between the TOTAL concentration of aliphatic VPH Components eluting within the range of interest and the total FID area count of the applicable VPH component peaks. A CF is calculated for C9-C10 aromatic hydrocarbons based upon a correlation between the concentration of the one aromatic VPH Component used to calibrate this range and the PID area count of this VPH component. |  |
|  | Is the %RSD of the calibration factor ≤20 for the surrogate and ≤25 for collective hydrocarbon ranges? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.4.2.12] |  |  | If the %RSD is ≤20 for Target VPH Analytes and the surrogate and ≤25 for hydrocarbon ranges, linearity can be assumed and the average CF can be used for quantitation in lieu of a calibration curve. |  |
|  | If collective hydrocarbon range concentrations are calibrated using a linear regression model, is the correlation coefficient ≥ 0.99? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 9.4.2.2 and 9.4.2.13] |  |  | Section 9.4.2.2: The use of CFs is the preferred approach to determine the relationship between the detector response and the Target VPH Analyte and hydrocarbon range concentrations. It is also permissible to utilize linear regression (see Sections 9.4.2.12 and 9.4.2.13). The linear regression approach for Target VPH Analytes and hydrocarbon ranges is described in Appendix 6. The use of non-linear regression is not allowed in this method and is considered a significant modification as discussed in Section 11.3.1. Section 9.4.2.13: In order for the linear regression model to be used for quantitative purposes, r (correlation coefficient) must be ≥0.99. |  |
|  | If using the linear regression model, are recalculated concentrations in the lowest standard using the final calibration curve within 70-130% recovery? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.4.2.13] |  |  | In addition, the resulting calibration curve from the linear regression must be verified by recalculating concentrations of the Target VPH Analytes and hydrocarbon ranges in the lowest calibration standard using the final calibration curve. Recoveries must be 70-130%. If recalculated concentrations from the lowest calibration standard are outside the 70- 130% recovery range, raise the RL to the concentration of the next highest calibration standard that exhibits acceptable recoveries when recalculated using the final calibration curve. |  |
|  | Is the initial calibration verified through the analysis of an initial calibration verification (ICV) made from a second source? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.4.2.15] |  |  | The initial calibration must be verified through the analysis of an initial calibration verification (ICV). This analysis must be performed every time an initial calibration is performed. The ICV must be prepared from a different stock standard than that used to prepare the calibration standard and must be analyzed immediately following the initial calibration. The ICV should be prepared at a mid-range calibration curve. |  |
|  | What is the acceptance criterion for the ICV? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.2] |  |  | The recoveries of all Target VPH Analytes and hydrocarbon ranges must be between 70- 130%. A new five-point calibration must be performed if >10% of all analytes are outside of criteria. |  |
|  | PROCEDURE – Sample Preparation | LAB | SOP | EXPLANATION |  |
|  | Does the laboratory use an autosampler or manual load system? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.2 Note] Answer: |  |  | Note: Although procedures for manual purge-and-trap load systems are provided below, MassDEP prefers the use of purge-and-trap autosamplers to reduce variability and to minimize the handling of samples for VPH analysis. If using a manual load system, see Section 9.1.2.1 for additional requirements. |  |
|  | Are samples introduced into the GC using a purge-and-trap concentrator? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.2] |  |  | Introduce volatile compounds into the GC using a purge-and-trap concentrator. |  |
|  | Is a 5.0 mL sample volume used? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.2.1] |  |  |  |  |
|  | Is the weight of the non-aqueous sample determined? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.3.1] |  |  | Weigh the sample vial to 0.1 g on a top-loading balance and determine the weight of the soil/sediment sample; this determination requires knowledge of the empty/tared weight of the sample vial and volume/weight of methanol preservative that was added to the sample vial. |  |
|  | Is the surrogate spiking solution added to non-aqueous samples prior to extraction? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.3.2] |  |  | Add a specified volume (recommended not to exceed 1.0 mL) of the surrogate spiking solution through the septum of the sample vial. The concentration and/or volume of the surrogate spiking solution may need to be increased for samples that are highly contaminated (based upon screening and/or field notes), to prevent dilution to below detectable limits. The amount of surrogate added should yield a final concentration of 2.5 mg/kg. |  |
|  | Is the sample agitated to mix? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.3.4] |  |  | Agitate sample to facilitate adequate mixing of spiking solution(s). |  |
|  | Is the sample allowed to settle until a layer of methanol is apparent? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.3.5] |  |  | Allow soil/sediment to settle until a layer of methanol is apparent. |  |
|  | For soil samples, is an appropriate aliquot of the methanol extract for sparging withdrawn through the septum of the sample container and added to 5 mL of reagent water for analysis? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 9.1.3.6 and 9.1.3.7] |  |  | Section 9.1.3.6: Using a microliter syringe, withdraw an appropriate aliquot of the methanol extract for sparging through the septum of the container. Sample screening data can be used to determine the volume of methanol extract to add to the 5 mL of reagent water for analysis. Section 9.1.3.7: Be advised that the volume of methanol aliquot added to the reagent water should not exceed 200 μL to preclude adverse solvent front and trap breakthrough difficulties. Alternatively, the addition of methanol extracts to reagent water can be performed in 40 mL VOC vials when an autosampler is used keeping similar methanol to water ratios. |  |
|  | Is the purge time 11 minutes? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.1.2.3] |  |  | Regardless if manual load or autosampler is used, purge the sample for 11 minutes. |  |
|  | PROCEDURE – Sample Analysis | LAB | SOP | EXPLANATION |  |
|  | When and how is manual integration performed? [NC WW/GW LCB Manual Integration Policy] |  |  | When manual integration is employed, the laboratory must clearly identify manually integrated compounds, document the reason the manual integration was performed, the date performed and who completed the work. A flag or qualifier code may suffice for simple manual integrations. In addition, a hardcopy printout of the data displaying the manual integration shall be included in the raw data package (i.e., both the original and manually integrated chromatograms, of similar scale, must be present in the data package). All information necessary for the historical reconstruction of data must be maintained by the lab. Additionally, the laboratory must employ a systematic data validation procedure to check manual integrations to assure integrations are technically sound and representative of the response. |  |
|  | Are collective peak areas for the hydrocarbon ranges integrated from baseline? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.5.4] |  |  | Collective peak area integration for the hydrocarbon ranges must be from baseline (i.e., must include the unresolved complex mixture "hump" areas). |  |
|  | Are individual target VPH analytes and surrogate compounds integrated using a valley-to-valley approach? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.5.4] |  |  | For the integration of individual target VPH analytes and surrogate compounds, a valley-to-valley approach should typically be used, though this approach may be modified on a case-by-case basis by an experienced analyst. In any case, the unresolved complex mixture “hump” areas must not be included in the integration of individual target VPH analytes and surrogate compounds. |  |
|  | Is the surrogate area excluded or subtracted from the collective hydrocarbon range areas? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.6.2] |  |  | When calculating the VPH by GC/PID/FID method aliphatic and aromatic hydrocarbon range concentrations, the laboratory must include the area of all peaks eluting within the retention time windows specified for these ranges, excluding surrogates. |  |
|  | QUALITY ASSURANCE | LAB | SOP | EXPLANATION |  |
|  | Is the Initial Demonstration of Laboratory Capability (IDLC) kept on file and updated when personnel are trained or when significant changes are made to the instrument or sample preparation? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.1.1] |  |  | Each laboratory that uses this method is required to operate a formal quality control program. The minimum requirements of this program consist of an Initial Demonstration of Laboratory Capability (IDLC) and an ongoing analysis of prepared QC samples to evaluate and document the quality of data. The laboratory must maintain records to document the quality of the data produced. |  |
|  | For the Initial Demonstration of Accuracy (IDA) and Initial Demonstration of Precision (IDP), does the laboratory prepare and analyze seven replicate Laboratory Control Samples (LCSs) according to the method? [MADEP VPH Rev. 2.1 (Feb 2018) Appendix 7 Section 3.0] |  |  | Prepare and analyze seven (7) replicate Laboratory Control Samples (LCSs) fortified at a concentration of 50% of the highest calibration curve standard (100 ug/L for aqueous samples and 5 mg/kg for soil/sediment samples). An LCS must be prepared and treated exactly as a typical field sample submitted for analysis, including exposure to all glassware, equipment, solvents and reagents. See Section 10.2.6 of the VPH by GC/PID/FID Method for how to prepare the LCS. |  |
|  | What are the acceptance criteria for the IDA and IDP? [MADEP VPH Rev. 2.1 (Feb 2018) Appendix 7 Sections 3.0 and 4.0] |  |  | IDA: The value derived for C (mean) must be within ± 30% of the true value or between 70 ug/L and 130 ug/L for aqueous samples and 3.5 mg/kg and 6.5 mg/kg for soil/sediment samples. IDP: Using the results calculated from Section 3.0 above, calculate the percent relative standard deviation (%RSD) of the seven (7) replicate analysis, as indicated below. The %RSD must be < 25 for both aqueous and soil/sediment samples. |  |
|  | Are mid-range Calibration Standards analyzed with each batch of up to 20 samples at the beginning and end of each analytical sequence? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 10.1.2, 3.5 and 9.4.3.2] List concentration: |  |  | Section 10.1.2: At a minimum, for each analytical batch (up to 20 samples of similar matrix), a beginning Initial Calibration or Opening mid-range Continuing Calibration Standard, closing mid-range Continuing Calibration Standard, LMB, LCS and LCSD must be analyzed. For analytical batches with more than 10 samples, the analysis of an additional mid-range Continuing Calibration Standard should also be considered. However, it should be noted that the analysis of the Continuing Calibration Standard is required prior to sample analysis, after every 20 samples, and at the end of an analytical sequence, at a minimum. Section 3.5: The continuing calibration standard is prepared from the same stock solution as calibration standards, and is generally one of the mid-level range calibration standard dilutions. Section 9.4.3.2: The concentration of the VPH Continuing Calibration Standard must be near the midpoint of the calibration curve. |  |
|  | What is the acceptance criterion for accuracy for the Continuing Calibration Standard? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.5.2] Answer: |  |  | For Target VPH Analytes and the surrogate, the %D or Percent Drift must be ≤ 20. For hydrocarbon ranges, the %D or Percent Drift must be ≤ 25. Greater %Ds or Percent Drifts are permissible for n-nonane (if included in the calibration of the C9 – C12 aliphatic range). If the %D or Percent Drift is > 30 for n-nonane, note the nonconformance in the laboratory narrative. |  |
|  | What corrective actions are taken if the Continuing Calibration Standard exceeds the acceptance criterion for accuracy? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.5.2] Answer: |  |  | If more than one Target VPH Analyte or hydrocarbon range fails to meet this criterion, the instrument must be recalibrated. Otherwise, sample analysis may proceed. |  |
|  | Are the Relative Retention Times verified and/or adjusted daily based on the Opening Continuing Calibration Standard? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 10.2.4 and 9.4.4] |  |  | Section 10.2.4: Relative Retention Times must be established for each Target VPH Analyte and hydrocarbon range of interest each time a new GC column is installed and must be verified and/or adjusted on a daily basis. (See Sections 9.3 and 9.4.4). Section 9.4.4: The range retention time windows must be established daily based upon the retention time of the marker compounds in the VPH Continuing Calibration Standard. |  |
|  | How are retention time windows determined for the analysis day? [MADEP VPH Rev. 2.1 (Feb 2018) Section 9.4.4] Answer: |  |  | Use the absolute retention time for each analyte in the continuing calibration standard as the midpoint of the window for that day. The daily retention time window equals the midpoint ± 3 times the standard deviation determined in Section 9.3. The marker compounds used for each range are defined in Table 5. |  |
|  | Is a Laboratory Control Standard (LCS) prepared and analyzed with each batch of up to 20 samples? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.1.2] |  |  | At a minimum, for each analytical batch (up to 20 samples of similar matrix), a beginning Initial Calibration or Opening mid-range Continuing Calibration Standard, closing mid-range Continuing Calibration Standard, LMB, LCS and LCSD must be analyzed. |  |
|  | What is the acceptance criterion for the accuracy of the LCS? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.6] Answer: |  |  | The spike recoveries for the Target VPH Analytes and the hydrocarbon ranges must be between 70% and 130%. |  |
|  | What corrective actions are taken if the LCS exceeds the acceptance criterion for accuracy? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.6] Answer: |  |  | If the recoveries are low and outside of the acceptance limits, reanalyze the LCS and associated samples. If still outside of the acceptance limits, recalibrate. If the recoveries are high and outside of the acceptance limits and the affected compound was detected in the associated samples, reanalyze the LCS and the associated samples. If recoveries are still outside of the acceptance limits, recalibrate. If recoveries are high and sample results were nondetect, data can be reported without qualification; however, the high recoveries should be noted in the laboratory narrative. |  |
|  | Is a Laboratory Control Standard Duplicate (LCSD) prepared and analyzed with each batch of up to 20 samples? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.1.2] |  |  | See above |  |
|  | What is the acceptance criterion for the precision of the LCSD? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.7] Answer: |  |  | The RPD for Target VPH Analytes and aliphatic and aromatic hydrocarbon range concentrations must be ≤ 25. |  |
|  | What corrective actions are taken if the LCSD exceeds the acceptance criterion for precision? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 10.2.7 and 10.2.6] Answer: |  |  | Section 10.2.7: See Section 10.2.6 for corrective actions associated with recoveries outside of acceptance limits. Section 10.2.6: If the recoveries are low and outside of the acceptance limits, reanalyze the LCS and associated samples. If still outside of the acceptance limits, recalibrate. If the recoveries are high and outside of the acceptance limits and the affected compound was detected in the associated samples, reanalyze the LCS and the associated samples. If recoveries are still outside of the acceptance limits, recalibrate. If recoveries are high and sample results were nondetect, data can be reported without qualification; however, the high recoveries should be noted in the laboratory narrative. |  |
|  | Is a Laboratory Method Blank (LMB) prepared and analyzed with each batch of up to 20 samples? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 10.1.2 and 10.2.3] |  |  | Section 10.1.2: At a minimum, for each analytical batch (up to 20 samples of similar matrix), a beginning Initial Calibration or Opening mid-range Continuing Calibration Standard, Closing mid-range Continuing Calibration Standard, LMB, LCS and LCSD must be analyzed. Section 10.2.3: A water or soil LMB is prepared by fortifying a reagent water blank (for aqueous samples), or 25 ml of methanol (for soil/sediment samples) with the surrogate spiking solution (using same volume of surrogate as samples). |  |
|  | Are the LMB and LCS analyzed prior to the samples? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.1.2] |  |  | The Initial Calibration or Continuing Calibration Standard, LMB, and LCS must be analyzed prior to samples. |  |
|  | What is the acceptance criterion for the LMB? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.3] Answer: |  |  | Peaks must not be detected above the RL within the retention time window of any analyte of interest. |  |
|  | What corrective actions are taken if the LMB exceeds the acceptance criteria? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.4] Answer: |  |  | If any of the performance standards specified in Section 10.2 are not met, the cause of the non-conformance must be identified and corrected before any additional samples may be analyzed. Any samples run between the QC samples that met the criteria and those that are fallen out must be reanalyzed, as noted in Section 10.2. These QC samples include the Continuing Calibration Standard, LMB, LCS, and LCSD. If this is not possible, the data must be reported as suspect. |  |
|  | What is the acceptance criterion for the surrogate recovery? [MADEP VPH Rev. 2.1 (Feb 2018) Section 10.2.8.1] Answer: |  |  | Each sample, LMB, LCS, LCSD, matrix spike, and matrix duplicate must be fortified with the surrogate spiking solution. Required surrogate recovery is 70% to 130% from each detector. |  |
|  | What corrective actions are taken if the surrogate recoveries exceed the acceptance criterion? Answer: |  |  | At a minimum, when surrogate recovery from a sample, blank, or QC sample is less than 70% or more than 130%, check calculations to locate possible errors, check the fortifying solution for degradation, and check for changes in instrument performance. If the cause cannot be determined, reanalyze the sample unless one of the following exceptions applies: - Obvious interference is present on the chromatogram (e.g., unresolved complex mixture); - Percent moisture of associated soil/sediment sample is > 25% and surrogate recovery is > 10%; or - The surrogate exhibits high recovery and associated target analytes or hydrocarbon ranges are not detected in sample. If a sample with a surrogate recovery outside of the acceptable range is not reanalyzed based on any of these aforementioned exceptions, this information must be noted on the data report form and discussed in the laboratory narrative. |  |
|  | Are soil and sediment results reported on a dry-weight basis? [MADEP VPH Rev. 2.1 (Feb 2018) Sections 9.1.4.1 and 9.6.3.] |  |  | Section 9.1.4.1: Soil and sediment results must be reported on a dry-weight basis. Transfer 5 to 10 g of sample into a tared (± 0.1 g) crucible. This sample must be obtained from a vial or container that does not contain methanol. Dry this 5 to 10 g sample overnight at 105°C and reweigh (± 0.1 g). Allow to cool in a desiccator before reweighing. Calculate the percent moisture of the sample using the equation provided in Section 9.6.3 (Equation 10). Section 9.6.3: Equation 10: Percent Moisture % Moisture= wet sample g-dry sample (g)wet sample (g)×100 Equation 11: Percent Solids % Dry Solids=100-% Moisture Equation 12: Dry Weight of Sample Wdg=(%Dry Solids)100×(g of extracted sample) |  |
|  | What is the laboratory’s reporting limit? [MADEP VPH Rev. 2.1 (Feb 2018) Section 12.0] Answer: |  |  | The RLs for hydrocarbon ranges shall be based upon the concentration of the lowest calibration standard for an individual analyte within the range of interest. The RL will be set at 100x the concentration of the lowest calibration standard for the associated analyte. Based on a concentration of 1 μg/L for the lowest calibration standard for all analytes, the following RLs would be generated for the hydrocarbon ranges: Aqueous Samples: Hydrocarbon range RLs would be equivalent to 100 μg/L. Soil/Sediment Samples: Hydrocarbon range RLs would be equivalent to 5 mg/kg based on a 1:1 ratio of methanol: soil and analysis of a 100 μL aliquot of the methanol extract in 5 mL water. |  |
|  | Is the data qualified on the Discharge Monitoring Report (DMR) or client report if Quality Control (QC) requirements are not met? [15A NCAC 02H .0805 (e) (5)] |  |  | Reported data associated with Quality Control failures, improper sample collection, holding time exceedances, or improper preservation shall be qualified as such. |  |

Additional Comments:

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

______________________________________________________________________________________________________________

Inspector: ______________________________________________________Date:____________________________________________